## Slide 1: MTS Testing Data Summary
Updated 2026-08-19

## Slide 2: Method Snapshot
Input: MTS 6DOF Book*.xlsx grouped as FE, LB, AR per intervention
Cycle basis: third cycle, centered avg(load/unload)
Model: M(theta) = C1*theta + C2*theta^2 + C3*theta^3
Plot orientation: angle on x-axis, moment on y-axis

## Slide 3: Intact

## Slide 4: Intervention: PUF Left

## Slide 5: Intervention: FUF Left

## Slide 6: Intervention: FBF

## Slide 7: Intervention: Posterior Release

## Slide 8: Intervention: SPO

## Slide 9: Flexion/Extension – All Interventions
similar stiffness up to ~8 deg
PLL release and SPO significantly decreased extension stiffness
Facetectomy decreased it as well but to a lesser extent
Continuous loss of stability with increasing bony/tissue resections
Flexion
Extension
For all interventions, stiffness is unchanged until around 8 deg
Facetectomy lowers stiffness afterwards
Extension
Flexion

## Slide 10: Lateral Bending – All Interventions
Left
Right

## Slide 11: Axial Rotation – All Interventions
Left
Right

## Slide 12: MTS Flexion: Comparison of Different Stiffness Metrics
Flexion
Extension

## Slide 13: MTS FE (Native Axes): Extension 0-5 and Flexion 0 to -5 Nm Secant Fit Panels

## Slide 14: MTS Flexion (Native Negative Axes): 0 to -5 Nm Secant Stiffness Bar

## Slide 15: MTS Extension (Native Positive Axes): 0 to 5 Nm Secant Stiffness Bar

## Slide 16: Comparison with Literature (Zhang 2020)

## Slide 17
Lamina Spreader Testing Data Summary

## Slide 18: Workflow Overview
1. Sync-check raw force and raw theta traces across interventions.
2. Apply force/theta filters only (no theta0 extraction yet).
3. Apply preload removal + theta0 baseline extraction + ramp-start detection with theta snap.
4. Mark transition-flagged cycles as ineligible.
5. Fit force-tip linear region using displacement >= 0 mm points (some cycles can fail fit criteria).
6. Fit moment-theta linear region using theta >= 0 deg points (some cycles can fail fit criteria).

## Slide 19: Step 1: Raw Force-Theta Sync Check
All interventions stacked before filtering/corrections.

## Slide 20: Step 2: Filtering Only (No Theta0 Yet)
Filtered force/theta signals only; no theta0 extraction or preload removal in this step.

## Slide 21: Step 3: Preload Removal + Theta0 + Ramp-Start Snap
Signals after preload removal, theta0 baseline extraction, and ramp-start detection with theta snap.

## Slide 22: Step 4: Mark And Exclude Ineligible Cycles
Step 4: Cycle Eligibility (Transition-Flag Filtering)
Intervention n_cycles eligible ineligible
---------------------------------------------------------
FBF 9 9 0
FUF 9 9 0
Intact 9 9 0
PUF 9 9 0
Posterior Release 10 10 0
SPO 9 6 3
Rule: transition-flagged cycles are excluded from stiffness fitting.
Note: ineligible here only counts transition_flag=True cycles.
Additional exclusions can occur later at fit stage (separate from transition flags):
Intervention force-fit excluded moment-fit excluded
--------------------------------------------------------------
FBF 0 0
FUF 0 0
Intact 1 1
PUF 0 0
Posterior Release 0 0
SPO 3 3

## Slide 23: Step 5A: Force-Tip Linear Region Fitting: Intact
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for Intact.

## Slide 24: Step 5A: Force-Tip Linear Region Fitting: PUF
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for PUF.

## Slide 25: Step 5A: Force-Tip Linear Region Fitting: FUF
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for FUF.

## Slide 26: Step 5A: Force-Tip Linear Region Fitting: FBF
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for FBF.

## Slide 27: Step 5A: Force-Tip Linear Region Fitting: Posterior Release
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for Posterior Release.

## Slide 28: Step 5A: Force-Tip Linear Region Fitting: SPO
Cycle-wise linear windows and per-cycle stiffness (fit uses displacement >= 0 mm; shared force y-limits) for SPO.

## Slide 29: Step 5B: Linear Stiffness Results Across Interventions
Mean ± SD from valid cycles (force-tip linear-region stiffness).

## Slide 30: Linear Stiffness Ranking
Step 6A Results: Linear-Region Force-Tip Stiffness
Rank Intervention median [N/mm] mean [N/mm] n_valid/n_cycles
--------------------------------------------------------------------------
1 SPO 8.158 7.037 6/9
2 FUF 8.158 7.851 9/9
3 Posterior Release 7.716 8.775 10/10
4 PUF 7.170 7.450 9/9
5 FBF 7.086 6.495 9/9
6 Intact 4.011 5.483 8/9

## Slide 31: Step 6A: Moment-Theta Processing: Intact
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for Intact.

## Slide 32: Step 6A: Moment-Theta Processing: PUF
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for PUF.

## Slide 33: Step 6A: Moment-Theta Processing: FUF
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for FUF.

## Slide 34: Step 6A: Moment-Theta Processing: FBF
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for FBF.

## Slide 35: Step 6A: Moment-Theta Processing: Posterior Release
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for Posterior Release.

## Slide 36: Step 6A: Moment-Theta Processing: SPO
Cycle-wise rotational stiffness processing (fit uses theta >= 0 deg; shared theta limits) for SPO.

## Slide 37: Step 6B: Rotational Stiffness Results Across Interventions
Mean ± SD from valid cycles (moment-theta stiffness).

## Slide 38: Rotational Stiffness Ranking
Step 6B Results: Moment-Theta Rotational Stiffness
Rank Intervention median [Nm/deg] mean [Nm/deg] n_valid/n_cycles
---------------------------------------------------------------------------
1 SPO 2.476 2.136 6/9
2 FUF 2.476 2.383 9/9
3 Posterior Release 2.342 2.663 10/10
4 PUF 2.176 2.261 9/9
5 FBF 2.151 1.971 9/9
6 Intact 1.217 1.664 8/9

## Slide 39: Appended Update: Fit-Point Rules
This version appends re-analysis results with explicit fit-point thresholds.
Displacement stiffness fitting: use only displacement >= 0 mm points.
Rotational stiffness fitting: use only theta >= 0 deg points.
All rankings/tables in this deck reflect these updated fit-selection rules.

## Slide 40: Appendix: Filters Used
Median filter (theta, kernel=5): replaces each sample with local median; removes isolated spikes while preserving step/ramp edges better than mean smoothing.
Butterworth low-pass (order=2): force cutoff=1.5 Hz, theta cutoff=1.2 Hz; attenuates high-frequency noise while preserving low-frequency cycle shape and phase reasonably well.
MAD despiking before low-pass: force window=11, z=3.5; theta window=15, z=3.5; robust outlier suppression using rolling median absolute deviation.
Pseudo-ramp rejection: keep cycles only if run_samples>=30, rise_time>=1.2 s, and theta_rise>=0.2 deg.
Duplicate-peak merge: close peaks are clustered and only strongest candidate retained (time-gap bounds 2.5 to 10.0 s).
Theta0 calibration filter: estimate zero-angle from low-force samples (<5% Fmax) but exclude low-force blocks with |median theta|>1.0 deg.
Ramp-start theta snap: move force-based ramp start to nearest theta~0 in early ramp window (1.5 s), unless start is already non-neutral (|theta|>1.5 deg).
Linear-fit point constraints: displacement fits use x>=0 mm; moment-theta fits use theta>=0 deg; flat low-slope segments removed before regression.

## Slide 41: Key Takeaways
Step-wise processing narrows from raw traces to quality-controlled cycles.
Results are reported in two complementary stiffness spaces:
force-tip (N/mm) and moment-theta (Nm/deg).
Rankings are based on median stiffness across valid cycles.
Fit-point constraints are displacement >= 0 mm and theta >= 0 deg.

## Slide 43: MTS Flexion/Extension – All Interventions
similar stiffness up to ~8 deg
PLL release and SPO significantly decreased extension stiffness
Surprising to see PLL release had a large impact
Facetectomy decreased it as well but to a lesser extent
Continuous loss of stability with increasing bony/tissue resections
Flexion
Extension
For all interventions, stiffness is unchanged until around 8 deg
Facetectomy lowers stiffness afterwards
Surprising to see PLL removal did not decrease stiffness
Extension
Flexion

## Slide 44: MTS FE (Native Axes): Extension 0-5 and Flexion 0 to -5 Nm Secant Fit Panels

## Slide 45: MTS Flexion: Comparison of Different Stiffness Metrics
Flexion
Extension

## Slide 46: Intervention Effect For Flexion

## Slide 47: Brainstorming – Model Scaling
Question 1: Does lamina spreader reliably characterize segment stiffness?
We’ll compare motion segment testing data for flexion between lamina spreader and traditional testing. If they are highly correlated, we can use spreader measurements as a proxy to joint stiffness. For example, we can then estimate motion segment stiffness using spreader.
Question 2: How do we quantify the impact of interventions on intact motion segments?
We’ll use traditional testing data to quantify the impact of interventions on intact motion segments. We’ll only use spreader if patient-specific assessment of the impact of interventions on stability is required. It’s technically not feasible to measure lateral bending and axial rotation using spreader.
Question 3: How do we estimate intact isolated motion segment stiffness from intact full spine patient stiffness?
We’ll measure/establish the relationship between these two features using lamina spreader. If successful, we can use intraoperative spreader measurements to estimate motion segment stiffness using either our cadaver measurements or the literature data.

## Slide 48: Cadaver Lab Goals
Goal 1: Establish usage guidelines:
  - Establish where to apply lamina spreader (left/right lamina, spinous process)
  - Establish how to perform consistent measurements (gradual force application, pause at the end, etc.)
  - Establish how many cycles are needed?
  - What makes a measurement cycle qualify?
Goal 2: Figure out whether to update stiffness on the go, or at the end of data collection
Goal 3: Collect data for multiple levels
- Figure out if there is a path to extrapolate multiple levels from one level measurement

## Slide 53
Summary for Intraoperative Stiffness Measurements Using Lamina Spreader

## Slide 54
Brown 2002 et al.

## Slide 55
Ebara 1992 et al.

## Slide 56: Comparison of Intraoperative Stiffness with Literature

## Slide 57: Comparison of Measurement Side

## Slide 58: Comparison of Spine Levels

## Slide 59
Standardized force band stiffness